Technical Note

Photon Counting with Low Repetition Rate Excitation

Peter Kapusta, PicoQuant GmbH, 2004

Question: I would like to use a N2 laser with 500 ps pulse width and 30 Hz repetition rate with my time-correlated single photon counting (TCSPC) system. Could you give me some suggestions how to combine the N2 laser and the TimeHarp 200?

Unfortunately, a 30 Hz repetition rate source cannot be used for TCSPC. There are practical limitations, that make time-correlated single photon counting impossible at such a low excitation rate. Due to the basic principle of the method, one can count photons only at 1-2 % of the excitation rate, whatever it is. With 30 Hz excitation rate, one must therefore limit the signal count rate to 0.3 Hz. Collecting enough counts to build up a decay histogram that can be analyzed at such a low count rate would at least take one day! Also, limiting the count rate to such a low value is everything but easy, because the dark noise rate of common detectors is in the order of 100–1000 Hz; even expensive cooled detectors feature 1–2 Hz. The above limitation is general and valid for all TCSPC boards, not only for the TimeHarp 200.

Question: OK, then what about to use a multichannel scaling (MCS) electronics, like the NanoHarp 250? It can process many photon counts per excitation cycle and in that way I can measure at higher signal rates.

No, not really. The average signal rate, i.e. your acquisition speed, will be still strongly limited due to one of the many pulse pile-up effects. Every single photon counting device has a pulse pair resolution limit which is related to the response speed of the detector. Simply, if the photon density exceeds a certain limit, the detector cannot respond to each photon with a corresponding single, well resolved voltage pulse. Two photons reaching the photocathode shortly after each other result in a merged double pulse at the detector output, as illustrated in fig. 1. As fast photon counting electronics react to a selected (usually the leading) pulse edge, such a double pulse is counted as one and the second photon is practically lost for the measurement, which has many consequences.

Fig. 1: If the photon density exceeds a certain limit, the pulse processing electronics cannot distinguish between two individual photons and the two photons are counted as one.
For example, the apparent count rate detected by the electronics is no longer a linear function of the light intensity. Therefore, detector suppliers sometimes provide a count rate correction curve, which shows the ratio of the hypothetical 'true' count rate and the actual output rate. The higher that ratio is, the more photons are lost. As an example, for the popular SPCM-AQR SPAD module, the correction factor is 2 at an output count rate of 9 MHz [1] and for the currently fastest detector, the microchannel plate photomultiplier, the same correction factor corresponds to 30 MHz output count rate [2]. However, there is no exact rate level that can be specified as a limit for pile-up free counting, because this factor converges to 1 asymptotically at low count rates. Though it is easy to calculate the true average count rate, this is not what we are after, as we want to record the unknown temporal evolution of the intensity. Loosing a photon means loosing timing information without a possibility of compensation and because intense signal periods are affected by pulse pile-up more than periods with low photon rates, the overall effect is a complex distortion of the histogram.

The effect is demonstrated in fig. 2 showing the count rate dependence of an instrument response function (IRF) measurement. The light source was a Xenon flash lamp with 200 Hz repetition rate. Rayleigh scattered 450 nm photons were detected with the PMA‑182 photomultiplier detector assembly and the signal was processed by the NanoHarp 250 multichannel scaler/photon counter. The output power of the flash lamp and the experimental runtime were kept constant, only the average detection count rate has been varied from 500 Hz to 30 kHz by gradual opening an iris in the emission detection arm.

Fig. 2: The effect of pulse pile-up demonstrated with an IRF measurement. As the detection count rate increases, the histogram becomes more and more flat, because the mostly affected channels are those around the peak, i.e. with the highest signal intensity.

A flattening of the decay curves recorded at higher detection rates is apparent. The region mostly affected by the pulse pile-up is around the peak intensity where the corresponding channels get less counts than would be proportional to the true temporal intensity. Due to the non-linearity of pulse pile-up distortion, further increase of the light intensity results in a dip in the histogram. In extreme cases, when the detector becomes saturated during the signal peak, the count number in the affected channels can even drop to zero.

The (in-)dependence of the histogram full width at half maximum (FWHM) on signal count rate can serve as a quick test for this kind of pulse pile-up distortion. Alternatively, the peak count value vs. the total counts (assuming constant acquisition time) can be plotted. Undistorted measurements would yield a constant FWHM and a linear plot of peak counts, respectively. In order to get meaningful results, the real signal shape must remain constant, only the amount of photons reaching the detector can be varied by means of an iris or neutral density filters. In practice, one approximates the signal rate by the total amount of registered counts divided by the acquisition time. Fig. 3 shows the evaluation of the histograms from fig.2.

Fig. 3: The dependence of the FWHM on the signal rate (left) and the nonlinear plot of peak counts (right) are manifestations of the pulse pile-up.

The meaning of the maximum average count rate
Let us assume, that we have a (nearly) ideal detector capable of responding to each photon properly up to fMAX = 10 MHz and the true photon rate never exceeds this limit (i.e. no photon will be lost). One might think that in this case the bottleneck will be elsewhere, but it is in fact not. Imagine the measured signal as a trail of flashes and consider the total length of time period when the photons are actually emitted (see fig. 4). During a τ [second] long flash, one can detect up to N = τ ·fMAX counts at best, which happens only when the photon rate in the flash is exactly fMAX [Hz] (a higher rate would lead to a counting loss). At an excitation repetition rate of fEXC [Hz], we have fEXC such flashes within one second and therefore fEXC · N counts per second. Hence, the maximum average count rate is RMAX = fEXC · τ · fMAX . The point is, that at low repetition rates (small fEXC) and short τ (short pulse or fast kinetics) the maximum average count rate RMAX ≪ fMAX, because all the photons are emitted only during a small fraction (= fEXC · τ ) of the total acquisition time. In principle, the limit of RMAX = fMAX can only be reached in steady-state measurements.

Fig.4: The average count rate equals to the total amount of detected photons within one second. But these photons are actually emitted during short time intervals at much higher rates.

To figure out how low RMAX is in real measurements, one can approximate τ by the FWHM of the IRF or by the average lifetime of the decay. Some typical measurement examples are given in table 1.

Measurement type (example):	RMAX for fmax = 10 MHz:
Xenon flash lamp, 200 Hz repetition rate, 500 ns pulse duration, IRF measurement	1000 Hz
N2 laser with 30 Hz frequency, luminescence decay measurement with τ = 20 ns	6 Hz
Flash lamp pumped Nd:YAG at 10 Hz, measurement of singlet oxygen decay with τ = 30 μs	3000 Hz

Table 1: Estimates of maximum average counting rates for some typical experimental conditions.

It should be noted that the values in table 1 are rough and optimistic estimates assuming constant photon rate during the pulse or decay which is obviously not the case. Furthermore, photon emission is a random process, therefore bursts with widely varying photon rates are reaching the detector. These statistical rate fluctuations were completely neglected as well as the dark noise background.

Conclusion

To exploit the sensitivity of time-resolved single photon counting, use an excitation repetition rate as high as your signal kinetics allow. In practice, fEXC is the only variable that can be altered when maximizing the RMAX = fEXC · τ · fMAX product; fMAX is ultimately limited by the detector and the effective τ is entirely determined by the photophysics of the sample.

References

[1] Data sheet for Single Photon Counting Module: SPCM-AQR Series,
http://optoelectronics.perkinelmer.com
[2]	Data sheet for Microchannel plate photomultiplier tube (MCP-PMTs) R3809U-50,	http://www.hamamatsu.com

PicoQuant GmbH
Unternehmen für optoelektronische Forschung und Entwicklung
Rudower Chaussee 29 (IGZ), 12489 Berlin, Germany
Telephone:	+49 / (0)30 / 6392 6560
Fax:	+49 / (0)30 / 6392 6561
e-mail:	photonics@pq.fta-berlin.de
WWW:	http://www.picoquant.com